# Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker*

Zhihao Jiang, Miroslav Pajic, Salar Moarref, Rajeev Alur,

and Rahul Mangharam

University of Pennsylvania, Philadelphia PA, USA

Abstract.  The design and implementation of software for medical de- vices is challenging due to their rapidly increasing functionality and the tight coupling of computation, control, and communication. The safety- critical nature and the lack of existing industry standards for veriﬁcation, make this an ideal domain for exploring applications of formal modeling and  analysis.  In  this  study, we  use  a  dual  chamber  implantable  pace- maker as a case study for modeling and veriﬁcation of control algorithms for medical devices in UPPAAL. We begin with detailed models of the pacemaker, based on the speciﬁcations and algorithm descriptions from Boston Scientiﬁc. We then deﬁne the state space of the closed-loop sys- tem based on its heart rate and developed a heart model which can non- deterministically  cover the whole  state space.  For veriﬁcation, we ﬁrst specify unsafe regions within the state space and verify the closed-loop system against corresponding safety requirements. As stronger assertions are attempted, the closed-loop unsafe state may result from healthy open- loop heart conditions. Such unsafe transitions are investigated with two clinical cases of Pacemaker Mediated Tachycardia and their correspond- ing correction algorithms in the pacemaker. Along with emerging tools for  code  generation  from  UPPAAL  models,  this  eﬀort  enables  model- driven design and certiﬁcation of software for medical devices.

Keywords:  Medical Devices, Implantable Pacemaker, Software Veriﬁ- cation, Cyber-Physical Systems.

## 1    Introduction

Over the past four decades, cardiac rhythm management devices such as pace- makers have expanded their role from “keeping the patient alive” to “making the patient’s life comfortable” . The addition of more safety and eﬃcacy features has resulted in  increased complexity, inevitably leading to  more safety violations. From 1996-2006, the percentage of software-related causes in medical device re- calls have grown from  10% to  21%  [1].  During the  ﬁrst half of 2010, the  US Food and Drug Administration (FDA) issued 23 recalls of defective devices, all

*  This research was partially supported by NSF research grants MRI 0923518,  CNS 0931239, CNS 1035715 and CCF 0915777.

C. Flanagan and B. Ko¨nig  (Eds.): TACAS 2012, LNCS 7214, pp.  188—203, 2012.

oc Springer-Verlag Berlin Heidelberg 2012

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        189

of which are categorized as  Class  I,  meaning there is  a  “reasonable probabil- ity that use of these products will cause serious adverse health consequences or death.” At least six of the recalls were caused by software defects [2]. Unlike other industries such as aviation and automotive, the safety concern in the medical device domain is focused on the physical plant, the patient in this case, rather than the controller. As a result, although in aviation and automotive industries, standards are enforced during software development, manufacturing, and post- market change [3,4], there are no well-established standards for development of software for medical devices. There is a pressing need for standards and tools to certify and verify the safety of software in medical devices. For device man- ufacturers, this has prompted recent interest in applying formal modeling and veriﬁcation techniques in medical devices software development [5,6].

In this eﬀort, we propose a Timed Automata representation of the heart and a dual chamber pacemaker. Our models and speciﬁcations are designed based on descriptions available from Boston Scientiﬁc [7,8], a leading manufacturer of pacemakers, and extensive medical literature on this topic. We then demonstrate how a model checker, like UPPAAL [9], can be used to ﬁnd safety violations and prove the correctness of medical device algorithms. We deﬁne the state space of the closed-loop system based on its heart rate. Unsafe regions can then be speciﬁed and the closed-loop system is veriﬁed against corresponding safety re- quirements. We also deﬁne unsafe transitions as the controller drives the open- loop plant from a safe state into an unsafe closed-loop state. We focus on two cases of unsafe transitions which are referred to as “Pacemaker Mediated Tachy- cardia (PMT)”. Modern pacemakers are equipped with correction algorithms to terminate these behaviors. We demonstrate how to identify known unsafe tran- sitions and prove the correctness of corresponding correction algorithms using model checker. The UPPAAL model developed in this paper is freely available online  [10].  These models can be used  as  a starting point for many purposes (e.g. to build models with costs and probabilities for quantitative analysis of the eﬃcacy of pacemaker algorithms; development of patient-speciﬁc algorithms). In particular, the veriﬁed pacemaker model can be automatically translated into Stateﬂow charts in Simulink for test generation and code generation [11].

The paper is organized as follows: In Section 2, we introduce the physiological and  timing  basics  of the  heart  and  pacemaker.  Section  3  presents  UPPAAL models  of the  basic  DDD  pacemaker  and  the  heart.  In  Section  4,  we  deﬁne unsafe regions and verify the basic pacemaker model against corresponding safety requirements. In Section 5, we proposed a procedure for identifying and verifying unsafe transitions and demonstrated using two cases of PMT.

## 2    Heart and Pacemaker Basics

The coordinated contraction of the heart is governed by its Electrical Conduction System (see Fig. 1). The Sinoatrial (SA) node, which is a collection of specialized tissue at the top of the right atrium, periodically spontaneously generates elec- trical pulses that can cause muscle contraction. The SA node is controlled by the

190         Z. Jiang et al.

nervous system and acts as the natural pacemaker of the heart. The electrical pulses ﬁrst cause both atria to contract, forcing the blood into the ventricles. The electrical conduction is then delayed at the Atrioventricular (AV) node, allowing the ventricles to ﬁll fully. Finally the fast-conducting His-Pukinje system spreads the electrical activation within both ventricles, causing simultaneous contraction of the ventricular muscles, and pumps the blood out of the heart.

Due to aging and/or diseases, the conduc-

Fig. 1. Cardiac electrical system

tion properties of heart tissue may change. These changes may cause timing anomalies in  heart  rhythm,  thus  decrease  the  blood pumping eﬃciency of the heart. These tim- ing anomalies are referred to as arrhythmias, and  are  categorized  into   Tachycardia   and Bradycardia.  Tachycardia  features  undesir- able fast heart rate which impairs hemody- namics. Bradycardia features slow heart rate which  results  in  insuﬃcient  blood  supply. Bradycardia maybe due to failure of impulse generation with anomalies in the SA node,

or failure of impulse propagation where the

conduction from atria to the ventricles is delayed or blocked.

Since the heart tissue can be activated by external electrical pulses, Brady- cardia can be treated by providing electrical pulses when the heart rate is low. Implantable Pacemakers  have been developed to deliver timely electrical pulses to the heart to maintain an appropriate heart rate and Atrial-Ventricular syn- chrony. Implantable pacemakers normally have two leads ﬁxed on the wall of the right atrium and the right ventricle respectively. Activation of local tissue is sensed by the leads, triggering Atrial Sense (AS) and Ventricular Sense (VS) events. Atrial Pacing (AP) and Ventricular Pacing (VP) are delivered if no sensed events occur within deadlines.

In order to deal with diﬀerent heart conditions, modern pacemakers are able to operate in diﬀerent modes. The modes are labeled using a three character system. The ﬁrst character describes the pacing locations, the second charac- ter describes the sensing locations, and the third character describes how the pacemaker software responds to sensing. In this work we describe the most com- monly used mode of pacemaker, the dual-chamber DDD mode that paces both the atrium and the ventricle, senses both chambers, and sensing can both activate or inhibit further pacing. Similarly, the VDI mode paces only in the ventricle, senses both chambers, and inhibits pacing if event is sensed. [12]

## 3    System Modeling

### 3.1    Timed Automata and UPPAAL

Timed automaton [13] is an extension of a ﬁnite automaton with a ﬁnite set of real-valued clocks. It has been used for modeling and verifying systems which are

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        191

triggered by events and have timing constraints between events. From the Boston Scientiﬁc pacemaker speciﬁcation [7], the pacemaker can be modeled using this Extended Timed Automata notation, which is a subset of formal semantics in UPPAAL. UPPAAL ( [9,14]) is a standard tool for modeling and veriﬁcation of real-time systems, based on networks of timed automata. The graphical and text- based interface makes modeling more intuitive. Requirements can be speciﬁed using Computational Tree Logic (CTL) [15] and violations can be visualized in the simulation environment.

### 3.2    System Overview

The function of a pacemaker is to manage the timing relationship between the atrial and ventricular events. Thus Timed Automata is  suitable for modeling both the deterministic behavior of a pacemaker and the non-deterministic behav- ior of the heart. The overview of the closed-loop system is showed in Fig. 2(a). The  heart  and  the  pacemaker  communicate  with  each  other  using  broadcast channels. The heart generates Aget!  and  Vget! actions, representing atrial and ventricular events that the pacemaker take as inputs. The pacemaker processes the  signals  and  generates  pacing  actions  AP!  and  VP!  to  the  corresponding components in the heart.

### 3.3    Basic DDD Pacemaker Modeling

The DDD pacemaker has 5 basic timing cycles triggered by events, as shown in Fig. 2(b). We  decomposed our pacemaker model into  5 components which correspond to the 5 counters. These components communicate with each other using broadcast channels and shared variables (as shown in Fig. 3).

Lower  Rate  Interval  (LRI):  This component keeps the heart rate above a minimum value. In DDD mode, the LRI component models the basic timing cycle which deﬁnes the longest interval between two ventricular events. The clock is reset when a ventricular event (VS, VP) is received. If no atrial event has been sensed  (AS), the component will deliver atrial pacing  (AP) after TLRI-TAVI. The UPPAAL design of LRI component is shown in Fig. 3(a).

Aget

AP

Heart

Pacemaker

Vget

VP

Fig. 2. (a) System Overview, (b) Basic 5 timing cycles of DDD pacemaker

VP!  >

AP

VP?     >

LRI

AVI

>

VS?     >

VP?              URI

>

(a) LRI component                      (b) AVI component                                 (c) URI component

Aget?

AS!>

>

VS?

AR>!

VP? >

PVARP

Vget?

>

VS!

>

VP?

>

VRP

AS?

VS?

VP?

ASVS??    >

(d) PVARP component                               (e) VRP component

Fig. 3. Components of the pacemaker model in UPPAAL

Atrio-Ventricular Interval (AVI) and Upper Rate Interval (URI):  The function of the AVI component is to maintain the appropriate delay between the atrial activation and the ventricular activation. It  deﬁnes the longest interval between an atrial event and a ventricular event. If no ventricular event has been sensed  (VS) within TAVI after an atrial event  (AS,  AP), the component will deliver ventricular pacing (VP). In order to prevent the pacemaker from pacing the ventricle too fast, a URI component uses a global clock clk to track the time after a ventricular event (VS, VP). The URI limits the ventricular pacing rate by enforcing a lower bound on the times between consecutive ventricle events. If the global clock value is less than TURI when the AVI component is about to deliver VP, AVI will hold VP and deliver it after the global clock reaches TURI. The UPPAAL design of AVI and URI component is shown in Fig. 3(b) and (c).

Post Ventricular Atrial Refractory Period (PVARP) and Post Ventric- ular Atrial  Blanking  (PVAB):  Not all atrial events (Aget!) are recognized as Atrial Sense  (AS!). After each ventricular event, there is a blanking period (PVAB) followed by a refractory period (PVARP) for the atrial events in order to  ﬁlter  noise.  Atrial events during  PVAB  are ignored and  atrial events dur- ing PVARP trigger AR! event which can be used in some advanced diagnostic algorithms. The UPPAAL design of PVARP component is shown in Fig. 3(d).

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        193

Ventricular Refractory Period  (VRP):  Correspondingly, the VRP follows each ventricular event (VP, VS) to ﬁlter noise and early events in the ventricular channel which could otherwise cause undesired pacemaker behavior. Fig. 3(e) shows the UPPAAL design of VRP component.

Parameter Selection:  Each timing parameter of the pacemaker has a feasible range. However, after those parameters are programmed, they are ﬁxed during pacemaker operation. Consider all possible combinations of feasible parameter values is infeasible. In this work, we only verify one instance of a DDD pacemaker with nominal values in clinical settings [8]. The values we choose are TAVI=150, TLRI=1000, TPVARP=100, TVRP=150, TURI=400, TPVAB=50.

### 3.4    Random Heart Model (RHM)

Aget

!

RHM-A

AP?                   Aget!

>

>

Fig. 4. RHM for the Atrial Channel

In  order  to  verify  pacemaker  algorithm, we need to ﬁrst deﬁne the state space for the  closed-loop  system.  The  state  space deﬁnition  should not  only  cover all  pos-

sible  pacemaker  operations,  but  also  be

physiologically  intuitive  for  safety  requirement  speciﬁcation.  To  this  end,  we deﬁne the state space of the closed-loop system by the atrial interval (interval between atrial events  ∈ {AS, AP}) and ventricular interval  (interval between ventricular events  ∈ {VS, VP}). This heart rate representation enables us to deﬁne unsafe regions for bradycardia and tachycardia.

The Random Heart Model  (RHM) is designed to cover open-loop heart be- haviors. It non-deterministically generates an intrinsic heart event Xget! within [Xminwait, Xmaxwait] after each intrinsic heart event Xget or pacing XP. Here we use two RHMs for the atrial and ventricular channel where X can be atrial (A) or ventricular  (V). RHM covers all possible input to the pacemaker if the interval  [Xminwait, Xmaxwait] is set to  [0, ∞]. It can also cover subset of pos- sible heart conditions by assigning appropriate values to those two parameters. The UPPAAL model of the atrial RHM is shown in Fig. 4.

## 4    Veriﬁcation Regarding Unsafe Regions

In this section, we deﬁne unsafe regions regarding bradycardia and tachycardia and specify two basic safety properties. These two basic safety properties are strict so that they must be satisﬁed by any pacemaker under all heart conditions. We then discuss reﬁnement of the safe regions and make stronger assertions.

### 4.1    Lower Rate Limit

The most essential function for the pacemaker is to treat bradycardia by main- taining the ventricular rate above a certain threshold. We deﬁne the region where the ventricular rate is slow, as unsafe. The monitor Pvv is designed to measure interval between ventricular events and is shown in Fig. 5(a). The property A[] (Pvv.two  a imply Pvv.t≤TLRI) is satisﬁed by the basic DDD pacemaker.

(a)                                               (b)

Fig. 5. (a) Monitor for LRL: Interval between two ventricular events should be less than TLRI, (b) Monitor for URL: Interval between a ventricular event and a VP should be longer than TURI

### 4.2    Upper Rate Limit

The pacemaker is not designed to treat tachycardia so it can only pace the heart to increase its rate and cannot slow it down. However, it is still important to guarantee it does not pace the ventricles beyond a maximum rate to ensure safe operation. To this eﬀect, an upper rate limit is speciﬁed such that the pacemaker can increase the ventricular rate up to this limit.

We require that a ventricle pace (VP) can only occur at least TURI after a ventricle event  (VS, VP). The monitor for the property is shown in Fig. 5(b) and the property A[] (PURI test.interval imply P URI test.t≥TURI) is satisﬁed by the basic DDD pacemaker model.

## 5    Veriﬁcation Regarding Unsafe Transitions

The two unsafe regions, introduced above, are intuitive but provide for loose safety  properties.  One  may  wonder  if we  can  further  reduce  the  safe  region. When  the  closed-loop  system  is  in  some  unsafe  state,  there  are two  possible scenarios. One is when, the open-loop plant without the controller, is  also in unsafe state. In our case, if the heart is in tachycardia, the pacemaker is not supposed to react so that this case is of little value to us. The other scenario is that the open-loop plant is in a safe state and the controller is driving the closed- loop system into some unsafe states. We call this scenario  Unsafe  Transition. In our case, the pacemaker may increase the heart rate inappropriately, which is referred to as Pacemaker Mediate Tachycardia (PMT).

We  now  introduce  two  cases  of  PMT  and  their  corresponding  correction algorithms. Since one closed-loop state may correspond to multiple execution traces, these PMT scenarios will not be returned by the model checker as counter- examples of safety requirements. However, we can still identify known PMT by adding constraints to the heart model or developing more complex requirements.

### 5.1    Veriﬁcation Procedure

The pacemaker manufacturers have developed anti-PMT algorithms to termi- nate diﬀerent PMT scenarios. In this section, we propose a general procedure to identify PMT scenarios and verify the safety and correctness of anti-PMT algorithms. The general steps for the procedure include:

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        195

SVT                    Bradycardia

AS ASASAS ASASASAS         AS                 AS

VS

VS         VS           VS              VS

0                        1000                2000                 3000                  4000        ms

(a)

PMT                  Appropriate

[AR]   [AR]  [AR]

AP

AP

AS      AS      AS      AS

I     I     I     I

l          l

vVP   VVP   VVP    VVP           VVS           V VS

b

0                        1000                2000                 3000                  4000        ms

(b)

Fig. 6. (a) SVT with ODO pacemaker (b) SVT with DDD pacemaker

1.  Show existence of PMT behaviors in the closed-loop system

2.  Introduce anti-PMT algorithms and check whether the two basic safety re- quirements still hold

3.  Prove correctness of anti-PMT algorithms by showing the non-existence of PMT scenarios

Here we use two well-identiﬁed PMT cases to demonstrate the methodology.

### 5.2    Veriﬁcation of the Mode-Switch Algorithm

Supraventricular  Tachycardia  (SVT):    SVT is an arrhythmia which fea- tures an abnormally fast atrial rate. Typically the AV node, which has a long refractory period, can ﬁlter most of the fast atrial activations during SVT thus the ventricular rate remains relatively normal. Fig. 6(a) demonstrates a pace- maker event trace during SVT, with a ODO mode pacemaker which just sensing in both channels. In this particular case, every 3 atrial events (AS) correspond to 1 ventricular event (VS) during SVT.

As an arrhythmia, SVT is still considered as a safe heart condition since the ventricles operate under normal rate can still maintain adequate cardiac output. However, the AVI component of a dual chamber pacemaker is equivalent to a virtual pathway in addition to the intrinsic conduction pathway between the atria and the ventricles. The pacemaker tries to maintain 1:1 A-V conduction and thus increases the  ventricular rate inappropriately. Fig. 6(b) shows the  pacemaker trace of the same SVT case with DDD pacemaker. Although half of the fast atrial events are ﬁltered by the PVARP period ([AR]s), the DDD pacemaker still drives the  closed-loop  system  into  2:1  A-V  conduction  with  faster  ventricular  rate, which is inappropriate. This problem can be resolved by switching pacemaker into single chamber mode to maintain appropriate ventricular rate.

Existence   of  PMT   during   SVT: Since PMT during SVT is an unsafe tran- sition,  we  need  to  ﬁrst  adjust  the  heart model  so  that  the  open-loop  behaviors covers SVT and are in the safe region. To this  end, the  interval  for the  ventricular RHM is set to [500,800]. This rate is slow enough not to be considered as tachycar- dia, but faster than the Lower Rate Limit

Fig. 7. Monitor  for  SVT:  Check  exis- tence of an endless sequence where the ventricular event interval ≤TURI

of the pacemaker so that pacemaker should not intervene. The monitor Pv  v is designed to show existence of PMT during SVT. It goes to the error state if the ventricular rate drops below the Upper Rate Limit (Fig. 7).

The existence property E[](notPv  v.err) is speciﬁed, which veriﬁes if there exists an execution in which the ventricular interval is always less or equal to TURI.  The  property  is  ﬁrst  veriﬁed  on  pacemaker  without  the  mode-switch algorithm. The property is satisﬁed during veriﬁcation.

Mode-Switch Algorithm: Intuitively, the mode-switch algorithm ﬁrst detects SVT. After conﬁrmed detection, it switches the pacemaker from a dual-chamber mode to a single-chamber mode. During the single-chamber mode, the A-V syn- chrony  function  of the  pacemaker  is  deactivated  thus  the  ventricular  rate  is decoupled from the fast atrial rate. After the algorithm determines the end of SVT, it will switch the pacemaker back to the dual chamber mode.

The mode-switch algorithm speciﬁcation we use is the same as the one used in Boston Scientiﬁc pacemakers  [8]. The algorithm ﬁrst measures the interval between  atrial  events  outside  the  blanking  period  (AS,  AR).  The  interval  is considered as fast  if it is above a threshold (Trigger Rate) and slow  otherwise (see Fig. 8  (1)). A  counter increments for fast  events and decrement for  slow events  (see Fig. 8  (2)). After the counter value reaches the  Entry  Count,  the algorithm will  start  a  Duration  which  is  a  time  interval  used to  conﬁrm the detection of SVT (see Fig. 8 (3)). In the Duration, the counter keeps counting. If the counter value is still positive after the Duration, the pacemaker will switch to the VDI mode (Fallback mode). In the VDI mode, the pacemaker only senses and paces the ventricle. At any time if the counter reaches zero, the Duration will terminate and the pacemaker is switched back to DDD mode.

In our UPPAAL model of the mode-switch algorithm, we use nominal param- eter values from the clinical setting. We deﬁne trigger rate at 170bpm (350ms), entry count at 8, duration for 8 ventricular events and fallback mode as VDI.

In order to model both DDD and VDI modes and the switching between them, we  made  modiﬁcations to  the  AVI  and  LRI  components. In  each  component

AS?>

Fast!

1

>

AR?

>

Interval   Slow!

AP?>

Fast?>

DDD!  >

Counter

>

du  beg!

VDI!

VS?

Duration

_       >

du_end!>

Slow?     C du_end?

VP?     D du_beg?

>

2

3

Fig. 8. Mode-Switch algorithm

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        197

two copies for both modes are modeled, and switch between each other when switching events (DDD, VDI) are received. During VDI mode, VP is delivered by the LRI component instead of the AVI component. The clock values are shared between both copies in order to preserve essential intervals even after switching. The modiﬁed AVI and LRI components are shown in Fig. 9.

AS?

VS?

VP?

DDD?

VDI?

AS?

VS?

AP

VP

LRI-

MS

AVI-

MS

AP?

DDD?

VDI?

Fig. 9. New LRI & AVI components

Veriﬁcation  Against  Basic  Safety  Requirements:    We  verify the  same basic safety requirements on the pacemaker model with mode-switch algorithm. The Upper Rate Limit property still holds but the Lower Rate Limit property is violated. When the pacemaker is switching from VDI mode to DDD mode, the responsibility to deliver VP switched from LRI component to AVI component. Since the clock reference is diﬀerent (Ventricular events in LRI and Atrial events in AVI), the clock value for delivering the next VP is not preserved. As a result, if an atrial event which triggered the mode-switch from VDI to DDD happens within [TLRI-TAVI, TLRI) after the last ventricular event, the next ventricular pacing will be delayed by at most TAVI time, which violates the Lower Rate Limit property (Fig. 11(a)).

Veriﬁcation of the  Algorithm:  We now present the veriﬁcation of the cor- rectness of the mode-switch algorithm by checking the same existence property E[] (not Pv  v.err) on pacemaker with mode-switch algorithm. We expect the vi- olation of this property, since during VDI mode the ventricular rate of the heart model is less than the Upper Rate Limit and will not trigger ventricular pacing. The counter example of the violation should show that mode-switch algorithm successfully switches the mode of the pacemaker to VDI mode. However, this property is still satisﬁed, indicating the mode-switch algorithm failed to elim- inate the  PMT scenario. Then we further restrict the  atrial interval of RHM to [100, 200]. Since the atrial rate for the new heart model is always above the trigger rate, mode switch to VDI mode should always eventually happen. The monitor PMS for the new property is shown in Fig. 10.

The property A<>  (PMS.err)  is not satisﬁed. The counter-example shows that some of the atrial events fall into the Post Ventricular Atrial Blanking period

Fig. 10. Monitor  for  Mode-Switch:  Check  if mode-switch  to  VDI  mode  will  always eventually happen

(PVAB) and got ignored. As a result, two fast  intervals may be considered as one slow interval (see Fig. 11(b)). If this happens more than one out of the Entry Count, mode-switch from DDD to VDI may never happen.

Discussion:  We demonstrated that model checking techniques can be used to identify unknown violations which cannot be identiﬁed during open-loop testing, showing  the  necessity  and  usefulness  of formal  veriﬁcation  in  medical  device software development and certiﬁcation. We also showed that adding new features to the veriﬁed system is a potential source for safety violations.

### 5.3    Veriﬁcation of Endless Loop Tachycardia  (ELT) algorithm

ELT overview:   The AVI component of a dual-chamber pacemaker introduces a virtual A-V pathway which forms a loop with the intrinsic A-V conduction pathway  (see Fig. 12(a)). In this scenario, a ventricular event  (VS) triggers a V-A conduction through the intrinsic pathway  (Marker  1 in Fig. 12(b)).  The pacemaker registers this signal as an Atrial Sense (AS) (Marker 2 in Fig. 12(b)). This event triggers VP after TAVI, as if the signal conducts through the virtual A-V pathway (Marker 3 in Fig. 12(b)). The VP will trigger another V-A con- duction and this VP-AS-VP-AS looping behavior will continue (see Fig. 12(b)). The interval between atrial events is TAVI plus the V-A conduction delay, which will drive the ventricular rate as high as the Upper Rate Limit.

From the pacemaker’s point of view, the pacemaker paces the ventricles as speciﬁed for every AS. That is why open-loop testing is unable to detect this closed-loop  behavior.  Modern  pacemakers  are  equipped  with  anti-ELT  algo- rithms to identify and terminate potential ELT. One common algorithm identi- ﬁes ELT by the ELT pattern and terminates ELT by increasing TPVARP time once to block the AS caused by the V-A conduction.

Existence of ELT:   As another case of unsafe transition, we again constrain the open-loop heart model into healthy heart. We set both the atrial interval

(a)                                               (b)

Fig. 11.  (a)  Safety Violation: VP is delayed due to the reset of timer during mode- switch,  (b) Correctness Violation: The blocking period may block some atrial events, turning two Fast events to one Slow event

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        199

Intrinsic

pathway

Fast “pathway”: pacemaker

A-V synchrony

(a)  Virtual  circuit  formed  by  the pacemaker and the heart

2

AS      AS        AS        AS      AS       AS

1                 3

VS       VP       VP       VP        VP       VP

ms

0                        1000                2000                 3000                  4000

(b) Pacemaker trace for ELT initialized by a early ventricular signal

Fig. 12. Endless Loop Tachycardia case study demonstrating the situation when the pacemaker drives the heart into an unsafe state [16]

and the ventricular interval above TURI so that ELT behavior is not covered by the heart model. Two monitors were designed to show the existence of ELT. One monitor, PELT det, shows the persistence of the VP-AS pattern and the other monitor, Pvv, shows that the ventricular rate is always no slower than the upper rate limit (Fig. 13). The existence property E[]  ((not PELT det.err) && (not Pvv.err)) fails on pacemaker without an anti-ELT algorithm.

The reason for the failure is that in our closed-loop system, AS can only be triggered by Aget signal from the  atrial heart model, where in  ELT case the AS is triggered by backward V-A conduction, which is not covered by our heart model.  In  order  to  solve  this  problem,  we  model  the  A-V  conduction  of the heart in addition to the orignal RHM. The adjusted RHM and the conduction component is shown in Fig. 14. For each atrial event Aget, the conduction com- ponent generates  V act  after  certain  delay  and vice versa. The  conduction is non-deterministic so that the old RHM is a special case for the new RHM. The PVARP and VRP components are also modiﬁed to accommodate new events A  act and  V act.

PELT_det                                                  Pv_v

Fig. 13. Monitor for ELT: VP-AS pattern detection and Upper Rate detection

After introducing the conduction component, the existence property holds, indicating the closed-loop system with new heart model covers ELT.

The ELT-termination Algorithm:  The ELT detection algorithm by Boston Scientiﬁc [7] utilizes these three features:

!

Aget

AP?                   Aget!

>

>

RHM-A

A_act!

|  | V |
| --- | --- |

Cond

!

get

VP?                   Vget!

V

>

V_act!

>

HM V

RHM-V

(a) Adjusted RHM            (b) New heart model                                    (c) Conduction component

Fig. 14. Modiﬁed heart model and the conduction component

—  Ventricular rate at Upper Rate Limit

—  VP-AS pattern

—  Fixed V-A conduction delay

The  pacemaker  ﬁrst  monitors  VP-AS  pattern  with  ventricular  rate  at  upper rate limit. Then it compares the VP-AS interval with previous intervals. ELT is conﬁrmed if the diﬀerence between the current VP-AS interval and the ﬁrst VP-AS interval are within ±32ms for 16 consecutive times. Then the pacemaker increases the PVARP period to 500ms once so that the next AS will be blocked and will not trigger a VP. ELT will then be terminated.

As the V-A conduction delays are patient-speciﬁc, the  algorithm compares VP-AS interval to a previously sensed value instead of an absolute value. Since we  can not store past  clock values  in  UPPAAL, we  can not explicitly  model this ELT detection algorithm. However, since the conduction delay in our heart model is within a known range, we can compare the VP-AS interval with this range. The VP-AS pattern detection module for our anti-ELT algorithm is shown in Fig. 15 (1). It detects the VP-AS pattern with ventricular rate at upper rate limit and sends out  VP AS event if the interval qualiﬁes.

A counter counts the number of qualiﬁed VP-AS patterns. It increases the PVARP  period  to  500ms  if  eight  consecutive  VP-AS  patterns  are  detected. (Fig. 15 (2)) The PVARP component is also modiﬁed so that the PVARP period can only be changed once by the anti-ELT algorithm. (Fig. 15 (3))

Veriﬁcation   Against   Bottom-Line   Safety   Requirements:     The  two bottom-line safety requirements still hold when the anti-ELT algorithm is in- troduced.

Veriﬁcation ofthe Algorithm:  The existence property E[]((not PELT det.err) && (not Pvv.err)) is not satisﬁed after the anti-ELT algorithm is introduced, in- dicating the algorithm successfully terminates ELT. We successfully reproduced the case when the algorithm works in the simulation environment of UPPAAL.

Discussion: In this case study, we showed that we may require the heart model to provide more physiological details when verifying more complex properties. We  also observed  some  limitations  of Timed  Automata when  modeling  more complex algorithms.

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        201

3

1

2

Fig. 15. Counter for VP-AS pattern

## 6    Related Work

Jee et. al present a safety assured development approach of real-time software using pacemaker as their case study in  [17]. They formally model and verify a single chamber VVI pacemaker using UPPAAL and then implement it and check the preservation of properties transferred from model to implementation code. Tuan et. al propose an RTS formal model for pacemaker and its environment and veriﬁed it against number of safety properties and timed constraints using PAT model checker [18]. They have modeled the pacemaker for all 18 operating modes as described in Boston scientiﬁc, but their work lacks speciﬁcation and analysis os complex behaviors of the pacemaker, such as mode-switch.

Wiggelinkhuizen uses mCRL2 and UPPAAL to formally model the pacemaker from  the  ﬁrmware  design  of Vitatron’s  DA+  pacemaker  [19].  Two  main  ap- proaches have been used to investigate the feasibility of applying formal model checking to the design of device ﬁrmware. The main approach consists of verify- ing the ﬁrmware model in context of a formal heart model and a formal model of a hardware module which fails for high heart rates because of the state explo- sion. Another approach is to verify a part of ﬁrmware design which was feasible and was able to detect a known deadlock rather soon.

Macedo et. al have developed a concurrent and distributed real-time model for a cardiac pacemaker through a pragmatic incremental approach. The models are  expressed  using  the  VDM  and  are  validated  primarily  by  scenario-based test, where test scenarios are deﬁned to model interesting situations such as the absence of input pulses [20]. The models cover 8 modes of pacemaker operation.

Gomes et. al present a formal speciﬁcation of pacemaker system using the Z notation in  [21]. They have also tried to validate that the formal speciﬁcation satisﬁes the informal requirements of Boston Scientiﬁc by using a theorem prover, ProofPower-Z. They have partially checked the consistency of their speciﬁcation through reasoning. No validation experiment regarding safety conditions were performed yet. [21]

Mery et. al in [22], formally model all operational modes of a single electrode pacemaker system using event-B and prove them. They use an incremental proof- based approach to reﬁne the basic abstract model of the system and add more functional and timing properties. They use the ProB tool to validate their models in diﬀerent situations such as absence of input pulses.

## 7    Conclusion and Future Work

In this paper, we modeled a dual-chamber pacemaker with advanced features us- ing Timed Automata. Timed automaton captures key features of the closed-loop system and enables the use of tools like UPPAAL in veriﬁcation. We then ver- iﬁed one instance of a dual chamber pacemaker model with nominal parameter values since it is impossible to consider all possible combinations. We deﬁned a heart rate representation of closed-loop state space and identiﬁed unsafe regions and unsafe transitions. We demonstrated that model checking techniques can be used to reveal safety violations which cannot be identiﬁed during open-loop testing. We also showed that adding features to previously veriﬁed system may result in  safety violations.  Furthermore, we  showed that  more  complex  heart model is need to provide more physiological insights during property speciﬁca- tion. The UPPAAL model developed in this paper is freely available online [10]. We hope that these models can be used as a starting point for many purposes (e.g. to build models with costs and probabilities for quantitative analysis).

In this paper, we only veriﬁed the safety and correctness of pacemaker algo- rithms. However, the ultimate goal for a pacemaker is to maintain the eﬃciency of the heart. As future work, we would like to evaluate the eﬃciency of those al- gorithms by assigning costs for diﬀerent heart conditions. The evaluation can be used to develop better treatment for general and speciﬁc patients. More complex heart models are therefore needed to provide physiological insights. However, rig- orous heart model reﬁnement should be considered to ensure model consistency. While Timed Automata is a good ﬁt for the problem studied here, it also has some drawbacks as it can not capture certain behaviors of some advanced algo- rithms like memorizing diﬀerence of clocks, and is also not scalable enough. Our future work will also focus on improving the eﬃciency of veriﬁcation toolchain for medical device certiﬁcation.

## References

[1]  List of Device Recalls, U.S. Food and Drug Admin. (last visited July 19, 2010)

[2]  Sandler, K., Ohrstrom, L., Moy, L., McVay, R.: Killed by Code: Software Trans- parency in Implantable Medical Devices. Software Freedom Law Center (2010)

[3]  AUTOSAR website: http://www.autosar.org/

[4]  AVSI website: http://www.avsi.aero

[5]  Alur,  R.,  Arney,  D.,  Gunter,  E.L.,  Lee,  I.,  Lee,  J.,  Nam,  W.,  Pearce,  F.,  Van Albert, S., Zhou, J.: Formal Speciﬁcations and Analysis of the Computer-Assisted Resuscitation  Algorithm  (CARA) Infusion  Pump  Control  System. Intl. Journal on Software Tools for Technology Transfer (STTT) 5, 308–319 (2004)

Modeling and Veriﬁcation of a Dual Chamber Implantable Pacemaker        203

[6]  ten Teije,  A., et  al.:  Improving medical protocols  by formal  methods. Artiﬁcial Intelligence in Medicine 36(3), 193–209 (2006)

[7]  PACEMAKER System Speciﬁcation. Boston Scientiﬁc (2007)

[8]  The Compass - Technical Guide to Boston Scientiﬁc Cardiac Rhythm Management Products (2007)

[9]  Larsen, K.G., Pettersson, P., Yi, W.: Uppaal in a Nutshell. International Journal on Software Tools for Technology Transfer (STTT), 134–152 (1997)

[10]  Jiang, Z., Pajic, M., Moarref, S., Alur, R., Mangharam, R.: Pacemaker UPPAAL model download: http://www.seas.upenn.edu/~zhihaoj/VHM/PM_verify.zip

[11]  Pajic,  M., Jiang, Z., Sokolsky, O., Lee, I., Mangharam, R.: From Veriﬁcation to Implementation: A Model Translation Tool and a Pacemaker Case Study. In: 18th IEEE Real-Time and Embedded Technology and Applications Symposium, IEEE RTAS (2012)

[12]  Barold, S., Stroobandt, R., Sinnaeve, A.: Cardiac Pacemakers Step by Step. Black- well Futura (2004)

[13]  Alur, R., Dill,  D.L.: A Theory  of Timed  Automata. Theoretical  Computer  Sci- ence 126, 183–235 (1994)

[14]  Behrmann,  G.,  David,  A.,  Larsen,  K.G.:  A  Tutorial  on  Uppaal.  In:  Bernardo, M., Corradini, F.  (eds.) SFM-RT 2004. LNCS, vol. 3185, pp. 200–236.  Springer, Heidelberg (2004)

[15]  Clarke, E.M., Allen Emerson, E.: Design and synthesis of synchronization skeletons using branching-time temporal logic. In: Logic of Programs, Workshop, pp. 52–71 (1982)

[16]  Jiang,  Z.,  Pajic,  M.,  Mangharam,  R.:  Model-based  Closed-loop  Testing  of  Im- plantable  Pacemakers.  In: ICCPS  2011:  ACM/IEEE  2nd  Intl.  Conf.  on  Cyber- Physical Systems (2011)

[17]  Jee,  E.,  Wang,  S.,  Kim,  J.K.,  Lee,  J.,  Sokolsky,  O.,  Lee,  I.:  A  Safety-Assured Development Approach for Real-Time Software. In: The Proceedings of 16th IEEE International Conference on Embedded and Real-Time Computing Systems and Applications, pp. 133–142 (2010)

[18]  Tuan, L.A., Zheng, M.C., Tho, Q.T.: Modeling and Veriﬁcation of Safety Critical Systems:  A  Case  Study  on  Pacemaker.  In:  Fourth  International  Conference  on Secure Software Integration and Reliability Improvement, pp. 23–32 (2010)

[19]  Wiggelinkhuizen, J.E.: Feasibility of Formal Model Checking in the Vitatron En- vironment. Master thesis, Eindhoven University of Technology (2007)

[20]  Macedo, H.D., Larsen, P.G., Fitzgerald, J.S.: Incremental Development of a Dis- tributed Real-Time Model of a Cardiac Pacing System Using VDM. In: Cuellar, J., Sere, K. (eds.) FM 2008. LNCS, vol. 5014, pp. 181–197. Springer, Heidelberg (2008)

[21]  Gomes, A.O., Oliveira, M.V.M.: Formal Speciﬁcation of a Cardiac Pacing System. In:  Cavalcanti, A., Dams,  D.R.  (eds.) FM  2009.  LNCS, vol.  5850,  pp. 692–707. Springer, Heidelberg (2009)

[22]  Mery, D.,  Singh, N.K.: Pacemaker’s  Functional Behaviors  in Event-B. Research report, INRIA (2009)